Supplementary Table 1.  
Organization and staffing of UPHS Center for Evidence-based Practice (CEP)

### History

CEP opened in July 2006 and is funded by the UPHS Chief Medical Officer.  CEP evidence reports look first at existing guidelines and systematic reviews, and review primary studies when the secondary data are not recent or complete enough to support informed decisionmaking. Local utilization and cost data are incorporated where appropriate, so reports can be tailored to local needs. CEP then works with key stakeholders to implement reports, including integrating them into computerized clinical decision support or quality initiatives, and measures their impact using administrative or clinical data.

### Staffing

- Director and co-director

Both trained in internal medicine, with additional degrees in epidemiology

Part of time spent in hospital practice

- Research analysts (2)

  - One with research doctorate (Ph.D) and 15 years experience in technology assessment

  - One with degree in health policy (M.S.) and experience in translation and dissemination of research findings into clinical practice.

- Center administrator

- Supporting personnel with primary appointments elsewhere in the medical center

  - Clinical liaison librarian (M.S.L.I.S)

  - Infectious disease specialist (M.D.) with additional degree in epidemiology

  - Outpatient internal medicine physician (M.D.) with extensive experience in medical editing and publishing

  - Pharmacist (Pharm. D.) with research doctorate (Ph.D) in health services research

- Close collaboration with the staff of the Department of Clinical Effectiveness and Quality Improvement.

- Consultations as necessary with other health system faculty and staff.

CEP staffing totals 4.5 full-time equivalents

### Activities

In 2009, CEP completed 24 full-length evidence reports, 15 evidence advisories, 2 clinical guidelines, and 1 economic study.  CEP also completed 7 manuscripts for submission to peer-reviewed journals, and worked on collaborative projects with the Centers for Disease Control and Prevention, the Healthcare Infection Control Practices Advisory Committee, the Association of American Medical Colleges, and the Society for Healthcare Epidemiology of America.  To promote evidence-based medicine, CEP offers academic detailing for hospital staff (physicians and nurses) and a training elective for residents.

CEP evidence reviews are shared publicly through the National Guideline Clearinghouse, the Cochrane indexed Centre for Reviews and Dissemination HTA Database, peer-reviewed publications, and conference abstracts.

Supplementary Table 2.  
System attributes considered in cardiac catheterization laboratory review

Shaded cells represent attributes where the specified device is less satisfactory than other devices or does not meet the desired standard.

Attributes of User Controls

| Attribute | Importance | Device “A” | Device “B” | Device “C” |
| --- | --- | --- | --- | --- |
| Control devices work in an intuitive manner | High | Yes | Yes | Yes |
| Control devices easy to distinguish from one another | High | Yes: size & shape | Yes: shape | No |
| Additional control devices | Low | Pistol grip control for C-arm and table | Second knob for table control | None |
| Zoom control | High | One step | One step | Multiple steps |
| Store and recall x-ray technics | Medium | Yes | Yes | Yes |
| Store and recall C-arm positions | Medium | 7 positions | 2 positions | 3 positions |
| One-touch C-arm park | Low | No | Push and hold | Push and hold |
| Voice recognition (optional) | Low | No | No | Yes |

Attributes of Patient Handling

| Attribute | Importance | Device “A” | Device “B” | Device “C” |
| --- | --- | --- | --- | --- |
| Pan from groin to heart without repositioning C-arm | High | Yes | † | Only with C-arm positioned at side |
| Table range of movement: head-toe (cm) | High | 110 | 100 | 125 |
| Table range of movement: left-right (cm) | Medium | 28 | 36 | 35 |
| Table range of movement: up-down (cm) | Medium | 78-108 | 76-104 | 78-110 |

†–Floor-mounted system can pan from groin to heart without repositioning C-arm.  Ceiling-mount system has not been tested in this respect

Attributes of Original Image Quality

| Attribute | Importance | Device “A” | Device “B” | Device “C” |
| --- | --- | --- | --- | --- |
| Can visualize Taxus stents with fluoroscopy technic | High | Most of the time | Most of the time | Sometimes |
| Quantum efficiency | High | 73%-78% | 75% | 75% |
| Spatial resolution (LP/mm) | High | ? | ? | 2.7 |
| Pixel size (µm) | High | 200 | 184 | 184 |
| Matrix size (pixels square) | High | 1,024 | 1,024 | 960 |
| Image size (cm square) | High | 20.5 | 18.8 | 17.6 |
| Choice of CRT or LCD display in exam room | Medium | Yes | No | No |
| Instant replay of cine images | Medium | No delay | No delay | ½ second delay |

Attributes of Review Image Quality

| Attribute | Importance | Device “A” | Device “B” | Device “C” |
| --- | --- | --- | --- | --- |
| Review and post-processing |
| Quality of images displayed on GEMnet | High | Good | Good | Poor |
| Images can be exported to CD-ROM | High | Yes, viewing software included on CD. Can omit identifying information | Yes, viewing software included on CD. Can omit identifying information. | Yes, viewing software included on CD. |
| Quality of images from CD-ROM displayed on Windows computer. | High | Good | Good | Good |
| Quality of images from CD-ROM displayed on Macintosh. | High | Cannot be displayed | Cannot be displayed | Cannot be displayed |

Attributes of Interoperability and Image Management

| Attribute | Importance | Device “A” | Device “B” | Device “C” |
| --- | --- | --- | --- | --- |
| Images display properly on GEMnet workstations still in service at PPMC. | High | Yes, reduced matrix size. | Yes, reduced matrix size. Quantitative analysis not always available for imported images | No: very noisy, brightness and contrast incorrect. |
| Images display at full resolution on GEMnet workstations | High | No (512 x 512) | No (512 x 512) | No (512 x 512) |
| DICOM-compliant | High | Yes | Yes | Yes |
| Can exchange data with department scheduling/ dictation computers | Medium | DICOM | DICOM | DICOM |
| Image storage on lab computer | High | 136,000 images | 100,000 images | 25,000 images |

Attributes of Quantitative Analysis

| Attribute | Importance | Device “A” | Device “B” | Device “C” |
| --- | --- | --- | --- | --- |
| User interface for quantitative tools | High | Point/click | Point/click | Console: Point/click but workstation interface less intuitive |
| User help | Medium | Context-sensitive | Context-sensitive | Context-sensitive |
| Automated segment outlining | Medium | Yes | Yes | Yes |
| Measures LV function | High | Yes | Yes | Yes |
| Measures arterial stenosis | High | Yes | Yes | Yes |
| Measures other distances | High | Yes | Yes | Yes |
| Calibration options | High | Catheter, subject-image distance, reference object | Catheter, subject-image distance, reference object | Catheter, table-object distance, reference object |

Attributes of X-ray Specifications

| Attribute | Importance | Device “A” | Device “B” | Device “C” |
| --- | --- | --- | --- | --- |
| Maximum frames per second | Medium | 30 | 30 | 30 |
| X-ray power (kW) | Medium | 100 | 100 | 80 |
| Heat capacity (kW) | Medium | 3,200 | 2,500 | 3,200 |
| Rotational angiography | Low | 3D reconstruction | 3D reconstruction | 3D reconstruction |

Attributes of Patient and Operator Safety

| Attribute | Importance | Device “A” | Device “B” | Device “C” |
| --- | --- | --- | --- | --- |
| Dose-reduction technology | High | Yes | Yes | Yes |
| Real-time radiation dose estimate | Medium | On-screen | On-screen | On-screen |
| Alarm when fluoroscopy on | High | Audible tone and icon on display | Warning light and icon on display | Warning light and icon on display |
| Alarm when cine on | High | Audible alarm and icon on display | Audible alarm and icon on display | Audible alarm and icon on display |
| Shielding reaches all areas | High | Yes | † | † |
| Patient collision avoidance | Medium | Proximity sensors halt C-arm automatically | Proximity sensors halt C-arm automatically | Proximity sensors halt C-arm automatically |

†–Not tested on ceiling-mount system

Attributes of Familiarity

| Attribute | Importance | Device “A” | Device “B” | Device “C” |
| --- | --- | --- | --- | --- |
| Number of PPMC cardiologists presently using manufacturer’s current imaging system | High | 7/7 | 3/7 | 6/7 |
| Number of PPMC cardiologists presently using manufacturer’s current image management system | Medium | 0/7 | 0/7 | 0/7 |

NHS Survey Results

| Attribute | Device “A” | Device “B” | Device “C” |
| --- | --- | --- | --- |
| Respondents describing image quality as “very good.” (attribute pertained to quality of images as seen in the lab) | Not included | 7/10 | 1/11 |
| Respondents describing overall ease of use as “very good.” | Not included | 1/10 | 8/11 |
| Respondents describing overall reliability as “very good.” | Not included | 2/10 | 3/11 |
| Respondents describing range of C-arm movement as “very good.” | Not included | 5/10 | 5/11 |

Responses reported as number of “very good” responses/total number of responses