Local porosity measurement from scanning electron microscopy images in backscattered electrons mode

Supplementary materials

Loïc Sorbier, Hedwige Poncet, Vincent Lecocq, Guillaume Maillet, Maroua Moula, Vincent Le Corre

IFP Energies nouvelles, Rond-point de l'échangeur de Solaize, BP 3, 69360 Solaize, France

Corresponding author : Loïc Sorbier, Tel : (+33) 4 37 70 29 69, Fax : (+33) 4 37 70 27 45, loic.sorbier@ifpen.fr

# Uncertainty in ε

The porosity ε reads:

|  | ε=VrZbxηb-ηVrZbxηb-η+VbZrxη-ηr | ( 1 ) |
| --- | --- | --- |

Let xi, i∈1,⋯,N, be N statically independent variables with uncertainties uxi. Let fxi be a moderately non-linear function of the variable xi. If the uxi are small then uf the uncertainty in fxi reads:

|  | uf2≈i=1N∂f∂xi2uxi2 | ( 2 ) |
| --- | --- | --- |

From equation ( 1 ) we obtain:

|  | ∂ε∂Vr=VbZbxZrxηb-ηη-ηrVrZbxηb-η+VbZrxη-ηr2=ε2VbZrxVr2Zbxη-ηrηb-η | ( 3 ) |
| --- | --- | --- |
|  | ∂ε∂Vb=-VrZbxZrxηb-ηη-ηrVrZbxηb-η+VbZrxη-ηr2=-ε2ZrxVrZbxη-ηrηb-η | ( 4 ) |
|  | ∂ε∂Zrx=-VbVrZbxηb-ηη-ηrVrZbxηb-η+VbZrxη-ηr2=-ε2VbVrZbxη-ηrηb-η | ( 5 ) |
|  | ∂ε∂Zbx=VbVrZrxηb-ηη-ηrVrZbxηb-η+VbZrxη-ηr2=ε2VbZrxVrZbx2η-ηrηb-η | ( 6 ) |
|  | ∂ε∂ηb=VbVrZbxZrxη-ηrVrZbxηb-η+VbZrxη-ηr2=ε2VbZrxVrZbxη-ηrηb-η2 | ( 7 ) |
|  | ∂ε∂ηr=VbVrZbxZrxηb-ηVrZbxηb-η+VbZrxη-ηr2=ε2VbZrxVrZbx1ηb-η | ( 8 ) |
|  | ∂ε∂η=-VbVrZbxZrxηb-ηrVrZbxηb-η+VbZrxη-ηr2=-ε2VbZrxVrZbxηb-ηrηb-η2 | ( 9 ) |

From which we recover equation ( 9 ) of the paper :

|  | uε2ε2≈ε2VbZrxVrZbx2× η-ηrηb-η2uVb2Vb2+uVr2Vr2+uZbx2Zbx2+uZrx2Zrx2+1ηb-η4η-ηr2uηb2+ηb-η2uηr2+ηb-ηr2uη2 | ( 10 ) |
| --- | --- | --- |

Equation ( 10 ) is undefined for η=ηb (ε=0) so that uncertainty must be obtained by another way. Let f=1-ε be the volume fraction of matter. Obviously we have uε=uf and:

|  | f=VbZrxη-ηrVrZbxηb-η+VbZrxη-ηr | ( 11 ) |
| --- | --- | --- |

The same procedure of uncertainty propagation leads to:

|  | uf2f2≈f2VrZbxVbZrx2× ηb-ηη-ηr2uVb2Vb2+uVr2Vr2+uZbx2Zbx2+uZrx2Zrx2+1η-ηr4η-ηr2uηb2+ηb-η2uηr2+ηb-ηr2uη2 | ( 12 ) |
| --- | --- | --- |

Equation ( 12 ) allows to evaluate the uncertainty in ε close to ε=0. It is not defined for η=ηr (ε=1).

# Uncertainty in η

Let E be the expectation and Var the variance. Figure 1 presents the detection chain for the acquisition of backscattered electron signal on a single pixel (see (Reimer 2010) p196). During the detection time, a number np of impinging electron is set on the sample. np follow a Poisson distribution with Varnp=Enp=N. A fraction of this incoming electrons End=Varnd=ηN enter the detector and are converted into a signal x with End=GηN. The variance of x reads :

|  | Varx=G2Varnd+Ep2=G2End+Ep2 | ( 13 ) |
| --- | --- | --- |

Where Ep is an electronic noise independent of the number of detected electrons.

Figure 1 : Scheme of the detection chain. Pre-amp.: pre amplifier; Amp.: amplifier; ADC: analog to digital converter.

The signal y=Cx+B coming out the amplifier is submitted to an electronic noise Ea independent of x so that :

|  | Vary=VarCx+B=C2Varx=C2G2End+C2Ep2+Ea2 | ( 14 ) |
| --- | --- | --- |

The standard uncertainty uy in the signal y reads:

|  | uy2=Vary=C2G2End+C2Ep2+Ea2 | ( 15 ) |
| --- | --- | --- |

The analog to digital converter adds another contribution to the digital signal y so that its standard uncertainty uy reads:

|  | uy2=uy2+112=C2G2End+C2Ep2+Ea2+112=CGy-B+C2Ep2+Ea2+112=αy+β | ( 16 ) |
| --- | --- | --- |

# Optimal working distance

Figure 2 : Scheme of the detector geometry.

Figure 2 describes the geometry of the annular detector. The relations between the internal radius of the detector ri, its external radius ro, the working distance Wd and the collection angle range αi;αo read:

|  | tanαi=riWd | ( 17 ) |
| --- | --- | --- |
|  | tanαo=roWd | ( 18 ) |

The angular distribution of backscattered electrons is well approximated by a Lambert’s law:

|  | dηdΩ=ηπcosα | ( 19 ) |
| --- | --- | --- |

Hence the detected signal ηd may be approximated by :

|  | ηd=ηπαiαocosαdΩ=2ηαiαocosαsinαdα=ηαiαosin2αdα=η2cos2αi-cos2αo | ( 20 ) |
| --- | --- | --- |

Using the trigonometric identity cos2α=1-tan2α1+tan2α in Equation ( 19 ) in combination with Equations ( 17 ) and ( 18 ) we obtain :

|  | ηd=η21-ri2Wd21+ri2Wd2-1-ro2Wd21+ro2Wd2=ηWd2ro2-ri2Wd2+ri2Wd2+ro2 | ( 21 ) |
| --- | --- | --- |

Figure 3 : Theoretical influence of the working distance on the detected signal for detector radii given by the microscope manufacturer ri=2.5 mm et ro=6 mm. Dotted line indicates the optimal working distance Wd.

Figure 3 shows the influence of the working distance on the detected signal with internal and external radii given by the microscope manufacturer. There exists an optimal working distance where the detected signal is the higher. Close to this maximum, the detected signal is very weakly dependent on the precise value of the working distance. From Equation ( 21 ) we obtain:

|  | ∂ηd∂Wd=2ηWdri2ro2-Wd4Wd2+ri22Wd2+ro22 | ( 22 ) |
| --- | --- | --- |

Then the optimal working distance Wd is obtained when ∂ηd∂WdWd =0:

|  | Wd=riro | ( 23 ) |
| --- | --- | --- |

Thus a small variation of 0.3 mm around Wd≈3.87 mm leads to a variation of signal less than 0.6 %.

If experimental backscattering signal y is concerned Equation ( 21 ) has to be corrected with the contrast (α) and brightness (β) ajustements and with a deviation of the indicated working distance given by the microscope (Wd) from the true working distance (Wd-δ). Then Equation ( 21 ) transforms into:

|  | y= αηWd-δ2ro2-ri2Wd-δ2+ri2Wd-δ2+ro2+β | ( 24 ) |
| --- | --- | --- |

Hence optimal indicated working distance is given by Wd=riro+δ. Figure 4 shows the mean grey levels measured on a same zone containing resin and massive alumina with the same adjustments for contrast and brightness but with varying the working distance. The continuous curves are a simultaneous least square fit of Equation ( 24 ) on both resin and alumina with only α, β and δ as free parameters. Values of ri=2.5 mm and ro=6 mm are given by the microscope manufacturer. The value for η for resin and alumina is taken from the Monte-Carlo simulations. The agreement between data and theoretical model of Equation ( 23 ) is very satisfactory. A very poor fit is obtained if the parameter δ is omitted.

Optimized values of the parameters are given in Table 1. The value obtained for δ (406 µm) is not surprising taking into account the practical difficulty to precisely calibrate the working distance on a SEM. The discrepancies between data and model shown on the three lowest working distance for resin are due to the damage done to the resin after seven frames acquisition.

| (a) | (b) |
| --- | --- |

Figure 4 : Mean grey levels y on (a) PMMA resin and (b) massive alumina as function of the indicated working distance Wd. Symbols are experimental data and continuous lines are the results of the least square fit of Equation ( 24 ).

Table 1: Optimized parameters for least square fit of Equation ( 23 ) shown in Figure ( 4 )

| Parameter | Value |
| --- | --- |
| α | 4134 |
| β | -50.94 |
| δ (mm) | 0.4058 |

# High resolution SEM images

A section of impregnated extrudates of sample B3000 was milled and thin slices were cut at a 300 nm thickness by a RMC PowerTome PT PCZ ultramicrotome fitted with a diamond knife. Thin sections were deposited on an aluminum stub and observed with a FEI Nova NanoSEM in low vacuum mode (10 Pa H2O) at 5 kV with a GAD detector.

Figure 5 : Low resolution SEM image of thin slice of sample B3000 showing rare cavities (arrows) due to porosity not impregnated with resin.

Figure 5 shows a low resolution SEM image of a rare part of the B3000 where cracks are observed. Bubbles due to improper resin impregnation are sometimes observed in these cracks. Figure 6 is a high resolution SEM image of the same sample. No evidence of unfilled pores are observed.

Figure 6 : High resolution SEM image of thin slice of sample B3000 showing no cavities due to porosity not impregnated with resin.

References

Reimer, L. 2010. Scanning electron microscopy: Physics of image formation and microanalysis, Second revised and updated edition. Heidelberg: Springer.